Specification for Interfacing

Radiology Information Systems

with IBM’s MedSpeak/Radiology

Revision 2.1

(modified 01/26/98)

Healthcare Industry Solutions

1.	Introduction

1.1.	Purpose

This document is intended for the software application developer, and systems integrator wishing to interface Radiology Information Systems (RIS), with IBM MedSpeak/Radiology. Also included in this document will be a description of IBM MedSpeak/Radiology, and the reason for the RIS interface requirement.

The reader should have a basic understanding of the HL7 (Health Level 7) protocol, and may want to refer to the HL7 Standard Specification, (version 2.1 or higher) for details. Specifically, the following chapters will be helpful in understanding the basic HL7 concepts needed for this interface:

·	Introduction

·	Control/Query

·	Order Entry

To obtain copies of this specification, contact the Health Level Seven organization:

HL7

3300 Woolitenau Ave, suite 227

Ann Arbor, MI 48104

(313)677-7777

hq@hl7.win.net

1.2.	Document Control

For questions or comments regarding this document, please e-mail:

medspeak_radiology@vnet.ibm.com

2.	IBM Medspeak Radiology Workstation

2.1.	Overview

MedSpeak/Radiology is a software application that facilitates the creation of a radiology report (film observation, or reading).  The application provides three main capabilities:

1.	 Rapid transcription of  dictation to digitized text.

2.	 Basic word processing: corrections, formatting, printing, etc...

3.	 Work flow processing:

·	Receive orders from foreign RIS, providing demographic, and other related information for exams.

·	Apply radiologist’s observations to orders, creating a report.

·	Allow acceptance and electronic signature of report, and transfer of report back to RIS.

It is the third capability, work flow processing, that brings about the necessity for a MedSpeak RIS interface.

2.2.	Information shared between MedSpeak and RIS

As mentioned above, MedSpeak accepts certain information from the RIS, for every completed examination. This information (patient demographics. physicians notes, etc...) is then available to the radiologists prior to (or during) a reading.

2.2.1.	Data sent from RIS to MedSpeak

Based on requirements gathered from several hospitals and radiologists, the following data will be accepted from the RIS, and saved in the RDW’s local database, accessible to the appropriate radiologists:

·	Patient Information

Id

Name

DOB

Gender

Phone Number(s)

Social Security Number

·	Order control

order status

accession number

Ordering provider

·	Observation Information

exam code (Universal ID)

time exam was completed

relevant clinical information (order notes)

·	Reserved fields (institution specified)

2.2.2.

Data sent from MedSpeak to RIS

Once the report has been accepted and signed by the radiologist, MedSpeak will send the following information back to the RIS:

·	Patient Information

ID

·	Order control

primary accession number

associated accession numbers

order status (Preliminary or Final)

time of status change

physician ID of radiologist

·	Observation

body of report

impression section

·	reserved fields (institution specified)

3.	Interface Protocols

The interface specified in this document describes the protocol for sending and receiving data elements, and the protocol by which that data is to be delivered. The protocol required to to send and receive data with the IBM MedSpeak is HL7, a health care industry emerging standard. IBM MedSpeak also requires that a specific communications protocol be used: TCP/IP client/server sockets.

3.1.	HL7

The HL7 protocol is a simple messaging system, in which all data is transferred as ASCII data, not unlike a simple text file. Each transaction consists of a message unit (file), which consists of segments (lines, or rows), which consist of fields. The fields are formatted depending on their HL7 data type. Each message, segment and field are variable length, and use ASCII characters for delimiters. An example portion of a message is included below:

MSH|^~/&|Radiology|IBMRDW|Radiology||19960412|ORU||||||

PID|||0123456||Smith^James^T||19610122|M

The first segment is a message header (MSH). It contains the delimiter characters to be used in the message (^~/&|), the message type (ORU), and other message control information. The PID segment contains patient demographics, such as name and date of birth. Notice the format of the DOB. This is an HL7 data type called TS (time stamp). TS is defined as YYMMDDHHMM[SS][=/-ZZZZ]. There are numerous HL7 data types, ranging from tightly restricted format (such as TS) to very loose restrictions (e.g.. ST, which is any string data). Also notice that some fields contain no data. These are optional fields.

For a more detailed description of HL7, refer to the HL7 Standard Specification (Version 2.1 or higher).

3.2.	TCP/IP Socket Communication

3.2.1.	Client/Server

The IBM MedSpeak uses a program called an Agent to transmit and receive HL7 messages using a client/server model over TCP/IP sockets. A client is the process that sends a message, and the server accepts (and processes) the message. The IBM MedSpeak Agent sends completed reports from its client port to the RIS server. The RIS client sends new orders to MedSpeak Agent server port. Acknowledgements are sent back over the same port that the messages are received from. In the following diagram, RDW refers to Radiology Dictation Workstation, which is now know as MedSpeak/Radiology:

Medspeak Agent will attempt to connect to the RIS server immediately upon startup, and if the RIS is unavailable, the MedSpeak Agent will retry the connection indefinitely, at a predetermined time interval (e.g.. every 30 seconds). The Medspeak Agent can be configured to reach the server over any valid TCP/IP port. See MedSpeak Installation and Administration Guide for details.

The Medspeak Agent server port can be any valid TCP/IP port. Again, see MedSpeak Installation and Administration Guide for details on how to configure the interface.

1.1.2.	HL7 LLP

All HL7 messages are transmitted using HL7 Lower Level Protocol, which simply means that each message is preceded with a start block character and terminated with the end block string.

4.	Inbound Orders (RIS --> MedSpeak)

4.1.	Steps To Sending MedSpeak a Radiology Order

The following steps are required to send a radiology order to MedSpeak:

1.	connect to the MedSpeak Agent server port.

2.	send an order to MedSpeak

3.	receive an acknowledgment from MedSpeak

4.2.	Connecting to MedSpeak Agent Server Port

Medspeak Agent can be configured with any valid TCP/IP port number. The RIS should simply issue a TCP/IP connect to that port. See MedSpeak Installation and Administration Guide for more details.

4.3.	Sending Orders

Medspeak expects unsolicited orders from the RIS when an exam has been completed. The order is sent as an HL7 ORU message. Medspeak inspects the message, checking for required fields. If they are not present, the message will be rejected. Optional fields will also be extracted and saved in the RDW’s local storage. All other segments and fields are allowed (if valid HL7), but will be ignored.

For all HL7 segement definitions in this document, the table columns are denoted as follows:

SEQ  		sequence number (field number)

Length	maximum length of field

Type		HL7 data type

RQ		required or optional

RP		repeating field

Name		field description

4.3.1.	MSH - message header

| SEQ | Len | Type | RQ | RP | Name |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | ST | R |  | Field Separator |
| 2 | 4 | ST | R |  | Encoding characters |
| 9 | 7 | ID | R |  | Message Type (ORU) |
| 10 | 20 | ST | R |  | Message Control ID |

MedSpeak will expect SEQ 9, the message type, to always be ORU. This can be overridden in agentapp.ini using the following attributes:

InboundOrderMessage	= < message type> (defaults to ORU)

InboundResultMessage	= < message type>

InboundResultMessage indicates the message type for new orders, updates and cancellations. InboundResultMessage indicates the message type for orders containing results ONLY. If InboundResultMessage is specified, it MUST be different than the InboundOrderMessage type.

A trigger event code can be included as a component of the message type, but will be ignored by medspeak (eg. ORU^R01)

4.3.2.	PID - patient information

| SEQ | Len | Type | RQ | RP | Name |
| --- | --- | --- | --- | --- | --- |
| 3 | 16 | CK | R |  | Patient ID (MRN) |
| 5 | 48 | PN | R |  | Patient Name |
| 7 | 8 | TS |  |  | Date of Birth |
| 8 | 1 | ID |  |  | Sex (M or F) |
| 11 | 106 | AD |  |  | Patient Address |
| 13 | 40 | TN |  |  | Phone Number (Home) |
| 14 | 40 | TN |  |  | Phone Number (Business) |
| 19 | 16 | ST |  |  | Social Security Number - Patient |

4.3.3.	OBR - observation request

| SEQ | Len | Type | RQ | RP | Name |
| --- | --- | --- | --- | --- | --- |
| 3 | 75 | CM | R |  | Filler’s Order Number (Accession Number) |
| 4 | 200 | CE |  |  | Universal Service ID (exam code) |
| 7 | 26 | TS | R |  | Observation date/time |
| 13 | 300 | ST |  |  | Relevant Clinical Info (order notes) |
| 16 | 60 | CN |  |  | Ordering Provider (Clinician) |
| 18 | 400 | CM |  |  | placer field 1 (user field - five components, see below) |

The placer1 field will contain up to five components, each designating a site-specific user field. These user fields each have a maximum length of 128.  If the user field is longer than the specified maximum, the message will be rejected. See the Medspeak Installation and Administration Guide for more details about configuring user fields.

4.3.4.	ORC - order control (optional)

| SEQ | Len | Type | RQ | RP | Name |
| --- | --- | --- | --- | --- | --- |
| 1 | 2 | ID | R |  | order control |
| 5 | 2 | ID | R |  | order status |

This segment is optional. If not included, medspeak assumes that this is order is a new order, or an update to an existing order (which are treated the same). If included, MedSpeak will examine the codes as follows:

·	If a new or updated order is being sent MedSpeak, the order control ID should be “SC” and the order status should be any valid HL7 status other than CA or CM (see below).

·	If an order is being sent as a cancellation, the order control ID should be “CA” (cancel order request) and the order status should be “CA” (order was cancelled).

·	If results are being sent in the order, the order control ID should be “SC” (status change) and the order status “CM” (complete). Alternatively, results can be sent as a different message type (see MSH description). Note: If the results were produced on MedSpeak, and MedSpeak is configured to use addendums (DeleteAccepted=No in agentapp.ini), the report will NOT be deleted from the MedSpeak database. The transaction will be ignored.

4.3.5.	OBX – observation results

This message segment is optional. If included, the OBX segment will contain the order notes (clinical history, procedure info, etc…). Order notes may also be sent in SEQ 13 of the OBR segment.

| SEQ | Len | Type | RQ | RP | Name |
| --- | --- | --- | --- | --- | --- |
| 1 | 4 | SI | R |  | Set ID (sequence number) |
| 2 | 2 | ID | R |  | value type (FT) |
| 3 | 80 | CE | R |  | Observation Identifier |
| 4 | 200 | ST |  |  | Observation Sub-id |
| 5 | 64K | FT | R |  | observation value |

The interface agent can be configured to filter the inbound OBX observation ids. The AGENTAPP.INI file contains an entry OBXIDFilter that is used for this purpose. OBX segments containing the observation ids specified in this entry will be added to MedSpeak's order notes. Example:

OBXIDFilter = /HISTORY/ALLERGIES/

NOTE: the first character specified is interpreted by MedSpeak as the delimiter character. The same character MUST appear at the end of the string, or the last id will not be recognized.

4.4.	Receiving Acknowledgements

Medspeak will send an HL7 ACK message back to the RIS following each received ORU message. The ACK will be formatted as follows:

4.4.1.	MSH - message header

| SEQ | Len | Type | RQ | RP | Name |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | ST | R |  | Field Separator |
| 2 | 4 | ST | R |  | Encoding characters |
| 9 | 7 | ID |  |  | Message Type (ACK) |
| 10 | 20 | ST |  |  | Message Control ID |

4.4.2.	MSA - message acknowledgment

| SEQ | Len | Type | RQ | RP | Name |
| --- | --- | --- | --- | --- | --- |
| 1 | 2 | ID | R |  | Acknowledgement code (see below) |
| 2 | 20 | ST | R |  | Message control id (control id from ORU MSH) |
| 3 | 80 | ST |  |  | Text Message |

The Acknowledgment code will be as follows:

AA	Application Accept - MedSpeak has accepted the order

AE	Application Error - MedSpeak was unable to perform the transaction. The text message (SEQ 3) will contain the error description.

AR	Application Reject - MedSpeak has rejected the order. The text message (SEQ 3) 		will contain a reason for the rejection.

4.5.	Notes

Medspeak will accept duplicate orders as “replacements, or corrections. That is, MedSpeak will simply replace order that has an accession number (SEQ 3 in OBR) equal to an existing order in the database.

If your system is unable to handle AR or AE acknowledgement codes, The AGENTAPP.INI file should specify: AlwaysAccept = Yes. This will cause the MedSpeak Interface agent to log the rejection in the reject log file, but it will respond with an “AA” code in the HL7 MSA segment.

4.6.	Example

The following message may be sent by the RIS:

MSH|^~\&|||||||ORU|0123456

PID|||012-345||Doe^John^P||19610122|M|||1 Fairway Lane||||||||012-34-5678OBR|||9901|26345^Head,left view|||199604181202|||||||||1234^Welby

Medspeak would provide the following acknowledgement:

MSH|^~\&|||||||ACK|2345678

MSA|AA|0123456

5.	Outbound Reports (Medspeak --> RIS)

5.1.	Steps To Receiving a MedSpeak Report

The following steps are required to accept a radiology report from MedSpeak:

1.	accept a connection from MedSpeak Agent client port

2.	Wait for messages from the client

3.	Read and process the message

4.	send an acknowledgment back to MedSpeak Agent

5.2.	Accepting Client Connections

The RIS Server should accept the client connection over a standard TCP/IP port. Medspeak must be configured with the appropriate port number. See MedSpeak Installation and Administration Guide for more details.

5.3.	Waiting for MedSpeak Reports

Medspeak will send the RIS a report at the following times:

·	A report has been just been accepted and signed by the radiologist.

·	The radiologist has just saved a report, and wishes all preliminary reports be sent to the RIS.

·	Medspeak has just established a connection with the RIS server, and has one or more reports pending delivery.

5.4.	Reading MedSpeak Report

The report is sent from MedSpeak as an HL7 ORU message. The table columns are denoted as follows:

SEQ  		sequence number (field number)

Length	maximum length of field

Type		HL7 data type

RQ		required or optional

RP		repeating field

Name		field description

5.4.1.

MSH - message header

| SEQ | Len | Type | RQ | RP | Name |
| --- | --- | --- | --- | --- | --- |
| 1 | 1 | ST | R |  | Field Separator |
| 2 | 4 | ST | R |  | Encoding characters |
| 3 | 15 | ST |  |  | Sending Application |
| 4 | 20 | ST |  |  | Sending Facility |
| 5 | 15 | ST |  |  | Receiving Application |
| 6 | 30 | ST |  |  | Receiving Facility |
| 7 | 26 | TS |  |  | timestamp |
| 9 | 7 | ID |  |  | Message type |
| 10 | 20 | ST | R |  | Message Control ID |

·	The message type (SEQ 9) will, by default, contain only the first component, which is “ORU”. If a trigger event code is desired, the administrator must place an entry in the agentapp.ini file as follows:

TriggerEvent=R01

·	SEQs 3-6 (sending and receiving application/facility strings) can be set using the following entries in the agentapp.ini file:

SendingApplication = <string>		;example: Medspeak Dictation

SendingFacility = <string>		;example: Radiology Lab #1

ReceivingApplication = <string>

ReceivingFacility = <string>

5.4.2.	PID - patient information

| SEQ | Len | Type | RQ | RP | Name |
| --- | --- | --- | --- | --- | --- |
| 3 | 16 | CK | R |  | Patient ID (MRN) |
| 5 | 48 | PN |  |  | Patient Name |
| 19 | 16 | ST |  |  | SSN |

5.4.3.	ORC

| SEQ | Len | Type | RQ | RP | Name |
| --- | --- | --- | --- | --- | --- |
| 1 | 2 | ID | R |  | Order control |

the order control will be either “CN” or “RE”. This allows “combined results” to be sent back to the RIS. See the examples section below for usage.

5.4.4.

OBR - observation request

| SEQ | Len | Type | RQ | RP | Name |
| --- | --- | --- | --- | --- | --- |
| 3 | 75 | CM | R |  | Filler’s Order Number (Accession Number) |
| 4 | 200 | CE | R |  | Universal Service ID |
| 20 | 640* | CM |  |  | filler field 1 (user fields, see below) |
| 22 | 26 | TS | R |  | results/rpt status change (date/time) |
| 25 | 1 | ID | R | Y | result status (Preliminary, Final, or Addendum) |
| 32 | 60 | CN | R | Y | Principal result interpreter (radiologist) |

·	The status change time is the date/time the report was accepted and signed by the radiologist (or last saved if status is preliminary).

·	MedSpeak allows a resident and attending physician to produce a report. In these cases, the status and radiologist fields will be repeating.

·	The filler-1 field will contain up to five components, each designating a site-specific user field. These user fields each have a maximum length of 128. See the Medspeak Installation and Administration Guide for more details about configuring user fields. The user fields can also be sent as “result sections” in the OBX. See the following section (OBX – observation) for details.

·	The status field (SEQ 25) will, by default, be P, F, or A as described above. The preliminary status id can be configured differently. To do so, the administrator must place an entry in the agentapp.ini file as follows:

PreliminaryStatusCode=<x>		;example: R (Result not verified)

5.4.5.	OBX - observation

| SEQ | Len | Type | RQ | RP | Name |
| --- | --- | --- | --- | --- | --- |
| 1 | 4 | SI | R |  | Set ID (sequence number) |
| 2 | 2 | ID | R |  | value type (FT) |
| 3 | 80 | CE | R |  | Observation Identifier |
| 4 | 200 | ST |  |  | Observation Sub-id |
| 5 | 64K | FT | R |  | observation value |

·	The Observation Identifier will contain the universal service id with a modifier of either “BODY” or IMP” describing the record as being part of the body of the report, or part of the impression section.  These string values can be reconfigured by including the following lines in the agentapp.ini file:

OBXBodyIdentifier = <string>		;example: REPORT

OBXImpressionIdentifier = <string>	;example: IMPRESSION

·	The “user fields” data can be sent back to the RIS in the OBX as a “results section” similar to the BODY and IMP sections. For example, if user field #2 contained a diagnostic code that the RIS expects to see in the OBX with an Observation Identifier of “DIAG”, then the following entry would be made in the agentapp.ini file:

OBXUserField2 = DIAG		;syntax: OBXUserField<1-5> = <string>

·	The universal Service Id field will contain ONLY THE FIRST id when the report contains an observation for multiple exams (accession numbers).

·	The Interface Agent can be configured to generate a single OBX section for each section (BODY and IMP). Paragraphs will be delimited by the HL7 escape sequence “\.br\”. The Agent’s default behavior is to generate a separate OBX for each paragraph, so to generate the single OBX, the administrator must place the following entry in the agentapp.ini file:

OBX=standard

·	The interface agent can also be configured to contain one “line” of result text. This is done using the OBXMaxLength configuration option.  An example, to send 80 character lines (ideal for typical dumb terminal display), the following would be coded in the AGENTAPP.INI file:

OBXMaxLength = 80

Note: OBXMaxLength may NOT be specified with OBX=standard.

·	The observation value for a Final or Preliminary will be the report text. For an Addendum, the observation value will contain the addendum only.  The original report will not be resent.

5.5.	Acknowledging the Report

The RIS should acknowledge MedSpeak in the same manner that MedSpeak acknowledges the RIS when receiving an order. See the section describing this above.

5.6.	Example

The following is an example of a report for a single exam:

MSH|^~\&|||||||ORU|2345678

PID|||012-345

ORC|RE

OBR|||9901|201A2&BODY^PA & Lat. Chest ||||||||||||||||||199604201055|||F||||||||D12345

OBX|1|FT|201A2&BODY^PA & Lat. Chest||FINDINGS: Comparison, 03/01/91.

OBX|2|FT||201A2&BODY^PA & Lat. Chest||The patient is status post right mastectomy. The lungs are clear.

OBX|3|FT|201A2&IMP^PA & Lat. Chest||No active cardiopulmonary disease. No interval change since 03/01/91.

A “combined result” may also be sent. This is a single report for multiple exams, using the “standard” OBX, and containing two user fields (ICD9 code and radiologists e-mail address):

MSH|^~\&|||||||ORU|2345678

PID|||012-345

ORC|CN

OBR|||9901|201A2^PA & Lat. Chest |||||||||||||||741.9^joerad@nyu.edu |||199604201055|||F|||||||D12345

ORC|CN

OBR|||9902|26545^MRI Thoracic Spine |||||||||||||||741.9^joerad@nyu.edu|||199604201055|||F|||||||D12345

ORC|RE

OBR|||9903|26546^CT Abdomen|||||||||||||||741.9^joerad@nyu.edu |||199604201055|||F|||||||D12345

OBX|1|FT|201A2&BODY^PA & Lat. Chest ||Lungs are clear. No pleural effusions.\.br\No evidence of vertabral body collapse.

OBX|2|FT|201A2&IMP^PA & Lat. Chest ||Patient’s lungs, abdomen and spine are normal.

Notice the use of the CN ORC segments for the first two exams, followed by a RE ORC which has the actual observation OBX’s following it. Notice also that when a combined result is sent back to the RIS, only the first (primary) procedure is specified in the OBX, because the MedSpeak application does not support the differentiation of procedures within a single report. That is to say, the one observation applies unilaterally to all the exams included in the message.

NOTE: MedSpeak will send the OBRs in the same order that the radiologist included the exams when creating the report. This may be helpful if the receiving system expects a “primary” exam distinguished from the other associated exams.